RADIOACTIVE MATERIALS RECEIPT REPORT

| Receipt Date: |  |  |  |  | Receipt Time: |  |  |  |  |  | Carrier: |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Received From: |  |  |  |  |  |  |  | Received By: |  |  |  |  |  |
| Note: Packages containing radioactive material are to be opened only by a qualified hazmat employee. |  |  |  |  |  |  |  |  |  |  |  |  |  |
| RAM Type: Normal Form Special Form Reportable Quantity (RQ) Package Type: Type A Type B Industrial Package (IP) I II III Excepted Package – Instruments/Articles Low Specific Activity (LSA) I II III Excepted Package – Limited Quantity Surface Contaminated Object (SCO) I II III Package Label: Excepted LSA Radioactive White I Radioactive Yellow II Radioactive Yellow III |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Source Mfr. & Model No. |  |  | Serial No. |  |  | Isotope/Form/Mass |  |  |  | Est. Activity |  | Date of Est. Activity |  |

| ## External Radiation Monitoring Results | Unless the package is exempt from monitoring requirements, test its radiation levels ASAP after receipt & no later than 3 hrs. after receipt if during work hours or not later than 3 hrs. from the beginning of the next working day if received after working hrs. |
| --- | --- |

| Survey Meter Mfr. & Model: |  | SM Serial No.: |  |  | SM Calibration Date: |  |
| --- | --- | --- | --- | --- | --- | --- |

| □ | Radiation Response Check |  | □ | Battery Check |  | □ | Visual Check |  | Background reading: |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

| Highest reading at package surface: |  |  | - Notify FDOH BRC if > 200 mR/hr at surface |
| --- | --- | --- | --- |
| Highest reading at 1 meter: |  |  | - Notify FDOH BRC if > 10 mR/hr at 1 m |

Notify RSO of any abnormal radiation levels

| ## Contamination Monitoring Results |  |  | Unless the package is exempt from contamination testing requirements (special form or gas), test it ASAP after receipt & no later than 3 hrs. after receipt if during work hours or not later than 3 hrs. from the beginning of the next working day if received after working hrs. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Swipe No. | Location | LSC Results (dpm) |  | Swipe No. | Location | LSC Results (dpm) | Swipe No. | Location | LSC Results (dpm) |
| 1 |  |  |  | 3 |  |  | 5 |  |  |
| Swipe No. |  | LSC Results (dpm) |  | Swipe No. | Location | LSC Results (dpm) | Swipe No. | Location | LSC Results (dpm) |
| 2 |  |  |  | 4 |  |  | 6 |  |  |

U.S. DOT package contamination limit = 2,200 dpm/100 cm2 or 6,600 dpm/300 cm2

| #### SAT/UNSAT |  |  |
| --- | --- | --- |
|  | Shipping papers (Bill of Lading/Declaration for Dangerous Goods) |  |
|  | Emergency Response Information (ERI) |  |
|  | Packing list |  |
|  | Contamination swipe analysis results |  |
|  | No evidence of damage that could impair safe operation |  |

| Notes: |  |
| --- | --- |