TRITIUM BIOASSAY ANALYSIS REPORT

| Name: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Principal Investigator: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date last handled: |  |  |  |  |  |  |  |  |  |  | Time: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date sample collected: |  |  |  |  |  |  |  |  |  |  | Time: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Date sample analyzed: |  |  |  |  |  |  |  |  |  |  | Time: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Maximum amount of 3H handled at any one time: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | (mCi) |  |  |  |  |  |  |  |  |  |
| Total amount of 3H handled since last bioassay: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | (mCi) |  |  |  |  |  |  |  |  |  |
| Location of use: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Instrument Mfr. & Model: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Serial No.: |  |  |  |  |  |  |  |  |  |  |  |
| Unquenched H-3 Std. Model: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | SN: |  |  |  |  |  |  |  | Activity (dpm): |  |  |  |  |  |  |  |  |
| Efficiency (%) “K”: |  |  |  |  |  |  |  |  |  | Results (dpm): |  |  |  |  |  |  |  |  |  |  |  | Counting time (min): |  |  |  |  |  |  |  |  |  |  |
| NIST certification activity date range: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| # Unquenched Background Std. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Results (dpm): |  |  |  |  |  |  |  | Counting time (min): |  |  |  |  |  |  |  |  |  |  | Total Counts “Cb”: |  |  |  |  |  |  |  |  |  |  |  |  |  |
| NIST certification activity date range: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | to |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

MDC Determination (per NRC NUREG-1507)

Note: This should be determined before counting samples submitted for bioassay

= = _____ uCi/L

Where:	Cb - is the unquenched background standard total counts for a ten minute count

K  - is the unquenched 3H standard counting efficiency for the instrument used

44,400 - includes the sample and background counting time “T” (10 minutes), and the unit corrections for converting dpm to uCi (2.22E6), and 2.0 ml to 1.0 liter (.002)

|  | / |  | = |  | x | 2.252 E -4 | = |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample (cpm) | Efficiency Correction |  |  | Sample (dpm) |  | Unit Correction |  | Activity (uCi/L) |

Results

*Whenever the H-3 activity concentration exceeds the calculated MDC for the counting equipment, consult U.S. NRC Reg. Guide 8.32 and the RSO for required actions.

| Performed by: |  |  | Date: |  |
| --- | --- | --- | --- | --- |